## Slide 1: Bad Channel Tests
Alexander

## Slide 2: The issue
High leakage current after wire-bonding
At least one noisy channel
  when increasing HV and current, channel noise drops belo average ans neighbors start to increase
Pulling this channel restores moderate current

## Slide 3: Possible Defects ?
A simple pinhole?
  unclear why bonding the AC pad causes current increase
  strip still connected via Rpoly to bias ring
Spark below poly?
  could damage Si bulk and oxide; current flows to bias ring independent of AC pad contact
Discharge event in center of sensor on AC pad
  pinhole plus defect in/near implant (pinhole possible with V>180V over coupling oxide)

## Slide 4: What to measure?
Revive IR camera
  check bond position AND center position of affected channel
  will we see through the Sylgard?
Check on center of sensor:
  Rint to neighbors on DC pads
  Rpoly of affected channel plus Rpoly of neighbors
  Ileak on DC and AC pads
More to be discussed afterwards…

## Slide 6
DC
poly
oxide
n+
p
passivation
AC