## Slide 1
Arjan Droste, With many contributions from Marit Bogert, Sandra de Vries & Marchien Boonstra
Delft Measures: Assessment of citizen science rainfall observations

## Slide 2: The Delft Measures Citizen Science project
Citizen science project, focusing on urban climate in Delft, the Netherlands
Over 50 Citizen Scientists
“Adopt” a weather station to measure
Alecto WS-5500
Research opportunities:
  Weather observations: rain, temperature, humidity, soil moisture
  Heat stress, UHI
  Benefits of citizen input for hydrometeorology
DMR data availability as of December 2023

## Slide 3: Citizen Science in (urban) meteorology
PWS: valuable for urban weather data
But: just a data-pull, without citizen interactions
Station setup unknown, which can cause massive errors

## Slide 5: Can we quantify the effect of ‘bad’ setup?
We assume an “ideal” installation, but that’s often not possible
People are limited by space, personal choices, skill
What is the effect of station installation on the observed PWS rainfall?
What are the hindrances people have for PWS measurements?

## Slide 6: Research questions
How accurate and reliable are the precipitation observations of the Alecto WS-5500, and is it possible to quantify setup differences?
Notes: Only 1st and 3rd question discussed today

## Slide 7: Research questions
How does the monitoring and maintenance of owners of PWS influence data quality?
Notes: Only 1st and 3rd question discussed today

## Slide 9: Green Village experimental setup
“Ideal” setup comparison: 1,2,3
Tilted station: 3, 4
Roof station: 6
Blocked station: 5
Environment variation: 7,8
Reference stations: radar, manual gauge, MRR, Davis VP2
Notes: Station 1 is used as a reference for the others

## Slide 10
Next to wall: expected underestimation
Trend shifts

## Slide 11: Errors in tipping volume
Lower tipping volume
Overestimation of rainfall
Incomplete emptying of tipping bucket
Tilted station

## Slide 12: Alecto 1 vs. references
Overestimation Alecto, ~10% higher than KNMI (WMO) manual gauge
Good correlation with gauge and radar
Large differences within (established) precipitation measurement techniques!

## Slide 13: Relative bias as function of placement
(near wall)
(rooftop)
(tilted)

## Slide 14: Interaction with DMR participants
Survey completed in September 2023
Questions on maintenance and monitoring frequency
Photos from participants
Home visits to assist with the installation;
Complex installation process leads to missing data
Notes: Opportunity for metadata

## Slide 15: DMR Participants
Variation in setups and bias, not always explained by photographs.
Limited options for station placement: realistic conditions
How often do they check their station & data?

## Slide 16: DMR Participants

## Slide 17: DMR Participants: maintenance
Frequency of maintenance variable amongst participants; cleaning often done less than once a month
No clear relationship between maintenance, setup information and bias

## Slide 18: DMR Participants: data check

## Slide 19: The Green Village setup
Accurate rainfall accumulations compared to KNMI references, overestimation of 10%
General quality of measurements quite good!
Lower tipping volume due to incomplete emptying of tipping bucket
Substantial biases for tilted and obstructed stations
Lower correlation on 5 min resolution due to sampling errors
How accurate and reliable are the precipitation observations of the Alecto WS-5500, and is it possible to quantify setup differences?

## Slide 20: DMR Participants
Large variation in bias for different stations
Large variation in maintenance frequency
Large variation in station setups
No clear relationship between station bias, maintenance or location from available data
How does the monitoring and maintenance of owners of PWS influence data quality?

## Slide 21: To conclude
Station setup is a strong component of bias and error in PWS (rainfall) data
Station quality vs setup quality
How to deal with this?
Citizen input is very valuable to correct PWS data
  And to foster interest & participation in science!
Currently: make link to local urban morphology

## Slide 22: Thank you for your attention
Arjan Droste
a.m.droste@tudelft.nl

## Slide 23: Filtering procedure
Following Netatmo QC algorithm from De Vos et al. (2019)
High Influx (HI), Faulty Zero (FZ) & Station Outlier (SO) errors + Bias Correction
Filtering based on neighbouring stations

## Slide 24: Filtering procedure
Following Netatmo QC algorithm from De Vos et al. (2019)
High Influx (HI), Faulty Zero (FZ) & Station Outlier (SO) errors + Bias Correction
Filtering based on neighbouring stations